Project no: COOP-CT-2003-508203

AURORA

Contact-free Dynamical Volumetric Measurements of Lower Body with Functional Clinical and Diagnostic Capacity

Instrument: Cooperative Research

Thematic Priority

Final Activity Report

Period covered: from 01.05.2004 to 31.05.2006	Date of preparation:30/07/2026

Start date of project: 01.05.2004			Duration: 30 Months

Project coordinator name: 				Sven Mooshake

Project coordinator organisation name		DIERS International GmbH

Content:

A. Publishable summary

B: Table of Contents Midterm Report

1.	Section 1 – Project objectives and major achievements	6

1.1.	Objectives overview	6

1.1.1.	Technical Objectives	6

1.1.2.	Scientific Objectives	6

1.1.3.	Societal and Economic Objectives	7

2.	Section 2 - Work package progress of the period	8

2.1.	Deliverable list	18

2.2.	Milestone List	19

3.	Section 3 Consortium management	19

3.1. Summarize of the project	19

3.1.	Consortium management task	20

4.	Other issues	22

C: ANNEX

Publishable final activity report

Objectives and overview:

Back and lower body musculo-skeletal problems are among the most frequent reasons for visits to the doctor and following work absence. Typically, X-rays or invasive methods are necessary to identify pathological changes, and it is rather difficult to objectively measure and quantify dynamic functional parameters of the lower body, as little equipment is developed for this purpose.

The objective of the project Aurora was to develop a (relative) low cost diagnostic tool able to reconstruct and analyse movements and dynamics of the lower back and legs using anatomically correct models of the musculo-skeletal system. The system will carry out three dimensional surface measurements over time to provide a full volumetric analysis of movements and should meet the following requirements:

- Non-invasive and no radiation

- Dynamic, full volume analysis

- Fast, automotive and easy to use

- None or little preparation time

The method is based on laser-generated stripe projection and reconstruction using triangulation. In a former EU project (4D BodyScan) it was shown that dynamic analysis of the human back is possible with a single one-sided white-light based system (this system is currently being world-wide distributed by Diers). For the Aurora project however, a laser based system was selected due to the additional challenge of separating 4 different devices and due to the fact that we needed a bigger field of focused light (scan field).

Contractors:

Four small and medium-sized enterprises (SMEs) and three university research centres (RTDs) from 4 European countries worked together on this project:

- Diers International GmbH (DE): (manufacturer of orthopaedic medical equipment; optical 3D and 4D measurement systems) coordinator of the project.

- Velomat Messelektronik GmbH (DE): (measure and controlling electronics) synchronisation of different components, data transfer.

- Smarttech Sp.ZO.O. (PL): (3D measurement systems; fringe analysis; micro interferometers) responsible for calibration, visualisation and animation.

- R&G Rimkevicius and Gintautas Company (LT): (optics and laser technology) development and building of the scanning devices.

- Katholieke Universiteit Leuven / Dept. of Biomechanics (BE): (biomechanics; mechanobiology; surface analysis) defining and analysing biomechanical models

- Politechnika Warszawska / Inst. of Micromechanics and Photonics (PL): (photonics and mechatronics; machine vision; positioning systems) image analysis, extraction of features required for the biomechanical models

- Eberhard Karls-Universität Tuebingen / Dept. of Sport Medicine (DE): (clinical analysis; treatment) provides a medical team responsible for defining and evaluation of relevant clinical parameters and –results.

Scientific achievements:

Four synchronised detecting systems are placed at different angles to define the scanning range, covering a scan area of approx. 1,2m³. The person to be measured either walks through or does some predefined movements within the scanning area.

“A set of movements within a normal range for the lower body are and will be defined – for bending or twisting, for example. If body movements are found to lie outside a certain normalised range, abnormalities may be present.”

The device must be thought of as a supporting system – it will not replace other clinical diagnosis methods – as it is intended rather to indicate where diagnosis is needed than to automatically do clinical diagnostics. Apart from the new laser development, only standard components are used to ensure that the price is kept low.

Technical specifications and device overview:

A split beam laser system with the capacity to project a pattern of 120-140 lines was developed (safety class 1M or 2M; energy < 1 mW and therefore safe for clinical applications). The combined scan frequency is currently approx. 12 Hz with a fan angle of 60°. A synchronisation unit controls the 4 single units. The synchronisation signal also allows synchronisation of external devices like force sensors, EMG etc.

Each single unit is independently calibrated and may be used as a stand alone system for single view analysis. A multi view calibration of the complete device is done on location. By using special transportation boxes the system is transportable and can be installed in relative small rooms (minimum free floor space 2.8m x 2.8m).

Soft-and hardware tools for data acquisition communicate with and synchronise the 4 sets of cameras, lasers and PCs. After recording, 4 sets of dynamic point clouds are generated (stripe detection with successive 3D reconstruction). After reconstruction, the 4 views are combined in order to generate a full volume dynamic model of the legs. On this volume model, various features and parameters are detected and extracted for the clinical and biomechanical analysis: fix-points and areas related to the skeletal structure (femur, patella, tibia, ankle, etc). All kind of metric measurements can be done on the model (distances, volumes, circumferences, angles, etc).

Special attention was also given to develop compression algorithms, as the amount of collected data is huge (50-100 MB/sec; depend on the reconstructed volume). After lossless compression, the size was significantly reduced.

Socio-economic relevance and policy implications:

The system does not rely on X-rays or other types of potentially harmful radiation, implying that this method is especially applicable for repeated measurements (“follow up” studies), as well as measurements of big population groups (“screenings”). This may have a beneficial effect on, or – in the best case – lead to a better understanding of health problems related to the lower body.

The measuring procedure itself demands no hard-to-learn skills, as no markers or palpation of the patient is necessary (contact-free measurement); allowing non-technical and even non-medical persons to operate the system.

Dissemination and use:

There is a wide range of possible applications for the device. The medical, biomechanical and clinical applications are evident: It could be used to evaluate ergonomics and wear-out injuries by scanning people going about their daily routine or by doing sport (e.g. early detection of asymmetry between the left and right parts of the body, which top athletes often develop). More general screening procedures in sports halls or fitness clubs is another possibility. One important and possible use is the screening of children and adolescents for early stages of posture abnormalities and defects, including scoliosis and knee joint defects. The project team works closely with doctors and medical specialists in order to discover suitable clinical applications. Experience has also shown that after introducing a new device on the market, a whole range of new applications can be developed from the input of the user groups.

The coordinator Diers expects to add the device to its current product range. Both clinical and ergonomic studies are still necessary in order to evaluate the models and methods, as well as establish and adapt user interfaces and -handling for acceptance of the end-user groups (medical, clinical and research).

The other SMEs, which are not active in the medical field, will make use of the new product in other fields, such as robotics, sports or computer games. However, much wider applications for these anatomically correct, dynamic models are possible, such as applications in the clothing and fashion industry.

Conclusions:

After finishing the project a working device is available. This device is capable of scanning, reconstructing and analysing anatomically correct models of the lower body in movement. The big advantage of this device is its capability to reconstruct a complete, surface-based 360° model of the moving person. On this model, an objective quantification of a large set of metric, anatomical and dynamic parameters and features is possible to extract, analyse and visualise.

However, as was already known at the beginning of the project, the tasks were rather challenging: We planned to develop a device that could automatically analyse (a part of) the human body without expert interaction. Apart from the technical challenges (laser, synchronisation), a new biomechanical methodology had to be evaluated and tested. This created some delays in the original schedule, and lead to the result that not all of the models have yet been evaluated thoroughly enough. Therefore, as follow up to the project, a sequence of clinical studies will be initiated before the product is released for production and marketing.

Keywords:

Contact-free optical 3D and 4D body measurement; Automatic detection of anatomical landmarks; Reconstruction of skeletal models; Lower back and legs.

# Section 1 – Project objectives and major achievements

## Objectives overview

The objectives of the project were separated in a technical, scientific, societal and economical part. In this report the focus is on the technical and scientific parts. Major goal is the development of an analysing system for the diagnosis of lower leg extremities.

The state of the art systems for motion analysing are marker based systems. Markerless analysing systems are rare on the field of human motion analysing. One system presented at a biomechanics conference (ISB 2005) is just a laboratory solution and far away from a commercial usable system. Furthermore other pattern analysing systems are restricted to small scan fields and often they do not support clinical demands and mainly they only have capacity for static analysing.

### Technical Objectives

The technical objective was to develop a prototype system that is able to capture anatomical surface data of the human lower limb, within a cube of 1,2 x 1,2 x 1,2 m and with frequencies up to 30 Hz. Preferably, split laser beam technology (based on raster triangulation) will be used.

From the technical point of view major topic was the development of the optical hardware. R&G developed a one dimensional system for the projection of the laser lines. R&G has decided to use this technique because this seems to be the best solution to fulfil the demands in this project. Therefore a split laser beam technology was developed which can project 120 horizontal stripes. The needed synchronization unit to synchronize camera and laser was developed by Velomat and R&G.

### Scientific Objectives

The scientific objective is the development of musculo-skeletal models of the lower limb, based on body surface data, and integrating biomechanics knowledge from literature and own research. The models will allow diagnostic measurements of the leg axis, patella tracking, foot disorders and volume detection.

The existence of a diagnostic device as described here will permit that more people can be scanned and diagnosed and hence the awareness of the population of the importance of proper treatment of musculo-skeletal diseases will be raised.

To get data from real surface structures of the lower body the still existing 4D-imager system was adopted to the legs. Even if the scan window is too small for the whole leg length, important data could be recorded. Based on the point clouds, generated by the DRACO software developed by DIERS, the partners responsible for image analysing and biomechanical modelling could work with this. Based on this WUT developed new algorithms for the point recognition. For validation of the real surface landmarks and the landmarks detected by the software skin markers were used. The automatic tracing during static and dynamic conditions is possible. The calculated xyz-coordinates finally can be used for the establishing of kinetic models developed by KUL. The needed anatomical reference points were defined by UKT.

Besides the detection of anatomical landmarks the processing and storage of the data was an important fact. For this purpose WUT developed special software.

There is sometimes a discrepancy between scientific purposes looking for objective values and end users who want impressive visualization and animation modes. Therefore the visualization software should provide solution for both: animation and quantification. This was in the responsibility of Smarttech and KUL to satisfy the demands of both groups.

### Societal and Economic Objectives

Especially in the field of clinical diagnosis the system offers a fast and non invasive diagnostic tool for the lower body. Especially the non radiation technique will open a wide range of application. Compared to x-ray analysing the system should have the capacity for dynamic diagnosis. If the system can be produced to an acceptable price and the handling is easy than it is expected, that the consortium will have a good position to establish the system in the clinical world.

# Section 2 - Work package progress

##### Work package T1 (technical): Initial equipment

| Work package number: | 1 | Start date or starting event: month 1 |
| --- | --- | --- |
| Participant id: | DIERS | KULEUVEN |
| Person-months per participant: | 3 | 1 |

Objectives:

A simulated grabbing system, able to grab single sided surfaces, will run in selected laboratories.

Description of work and progress:

#### Task T1.1 Prepare and adapt existing equipment (DIERS, KULEUVEN)

The system available for scanning the back shape was adapted to the recording of the legs. The system was located at UKT. Using this system important pre work was possible to provide the other partners with needed information about surface shape and first data from stripe images to start the development of the image analysing algorithms.

##### Work package T2 (technical): Prototype development

| Work package number: | 2 | Start date or starting event: month 1 |  |  |
| --- | --- | --- | --- | --- |
| Participant id: | DIERS | R&G | VELOMAT | SMARTTECH |
| Person-months per participant: | 24 | 48 | 36 | 6 |

Objectives:

Develop/assemble 4-device scanning equipment, based on laser multi-beam raster stereography.

Description of work progress:

#### Task T2.1	Develop optical and laser components (R&G)

#### Task T2.2	Integrate optical and laser components, including cameras (R&G, VELOMAT)

The produced device allowed frame acquisition rates of up to 40-60 Hz, depending mainly on the number of lines that should be projected on the object. In order to achieve maximum frame rate of the scan the software allowed successive scans alternately in forward and backward directions; thus giving a double frame rate. Nevertheless, the preliminary results indicate, that for the generation of about 120 lines at a fan angle of about 60 degrees the scan speed limit is below 60 Hz. The faster frame rates can be accomplished only by reducing the number of the lines in the pattern (e. g. at the maximal 30 Hz scanning rate about 220 lines can be generated, while at the maximal 60 Hz scanning rate only 120 lines can be projected on the object). Reduction of the distance between the adjacent lines in the object plane (angle between the lines) also can increase the scanning speed, though not so strongly: the reduction of the fan angle of the generated lines from 60 to 30 degrees can increase the maximum scanning frame rate only by 10-20%. Therefore the reduction of the angle between the lines, although possibly helpful in certain circumstances, in general can not increase the scanning speed significantly.

However, not all benefits of the optical set-up proposed were exploited during the project. Coherent and dynamic properties of galvanometer scanning projector system using straight optics, which might lead to quite specific development of such volumetric systems like developed during the project, left behind future development and our pre-existent knowledge. That is not actually drawback of work of current working team, as it was difficult to foresee such development possibility when starting from experience with static white light systems used before.

Compared to the first prototype some improvements in laser intensity were done at the second version. Furthermore differences in light intensity between the four lasers were diminished.  Using new versions of processors and improved firmware to control the lasers, this problem could be solved.

First measurements in Tübingen revealed that, even with the increased laser intensity, the recorded images were not good enough for analysing. Tests with additional illumination failed because of problems with shadowing and blending. Due to the fact that an increase of the laser intensity is not possible with the available system the cameras from Kamera Werke Dresden must be changed. The new camera used in the system is a monochrome 1024x768 pixels camera from POINT GREY Research Inc. called FLEA. Although it can be used with the same SDK some adjustments in grabbing software had to be made. These cameras are offering better sensitivity to get good image quality from the laser stripes.

Task T2.3	Integrate force sensor components (VELOMAT)

#### Task T2.4	Implement control and electronic interface software (VELOMAT)

#### Task T2.5	Implement data channel interfaces (DIERS, VELOMAT)

The development of a foot sensor board to analyse pressure distribution was done by VELOMAT. They developed the sensor plate SPL-0030 for pressure distribution analyses. The dimension of the sensor area is 250x320mm² without gaps. The sensor resolution for the area is 5x5mm². The special electronic development makes it possible to combine several single plates to one plate with an increased measuring area. It is possible to arrange the board in different orders and get sizes for example of 500x320mm² (2 board’s) or 1000x640mm² (4 board’s). This will offer a good flexibility for the use in different clinical applications.

A second plate was constructed to analyse the centre of pressure (COP) during standing. This based on the integration of 4 sensors mounted at the 4 corners of the standing platform. Using special software the trajectories of the centre of gravity and the centre of pressure can be displayed.

As the synchronization unit serves as the basic unit for the synchronization of camera and laser beam this signal also serves for the synchronization of other signal sources. In agreement with the requirements for channel synchronization, formulated by UKT, the master synchronization unit has the possibility to send or to receive an analogous signal to external measurement devices.

#### Task T2.6	Calibration (DIERS, SMARTECH)

Single calibration

SMARTTECH took an effort to reanalyze the problem of single calibration. In its previous version, a calibration was based on simplified mathematical model which was causing relatively high level of systematic errors. Main sources of this error came from the fact that real image forming principle for stripe projector and detector pair is superposition of different phenomenon’s like lens distortion or perspective. The current calibration scheme is based entirely on experimental assessment of calibration function. There has been developed a whole bunch of different sophisticated interpolation and extrapolation algorithms over custom data structures that contain 2D and 3D data. The result of this development is a significant drop of error – magnitude decreased about 5x. Main efforts were done for increasing an accuracy of single system calibration. Quality of this calibration is important and basic factor for further quality of whole set of features including multiple calibration, landmark detection and cloud skinning

Multi View Calibration

The entire method of multicalibration was changed within last 6 months. The need for evaluating a new method arose from negative results of evaluation multiple calibration accuracy. Also great difficulties were met during process of calibration with use of previous method.

To remind, the old calibration was based on measuring 3 positions of planes for each pair of systems. The new calibration is based on sphere measurement and consists of measuring 3 positions of sphere. The first great advantage of new method is that it provides user with stable error range, which is easy to refine by simply making additional sphere measurement. Secondly, it involves practically no hardware except one sphere-shaped ball. For current setup of systems, we provided such ball which was measured and verified for out-of-sphere errors.

##### Work package T3 (technical): Data grabbing and synchronisation

| Work package number: | 3 | Start date or starting event: month 6 |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Participant id: | VELOMAT | DIERS | UKT |  |  |
| Person-months per participant: | 12 | 6 | 1 |  |  |

Objectives:

Data grabbing and synchronisation

Description of work progress:

#### Task T3.1	Hardware synchronisation of data channels (DIERS, VELOMAT)

#### Task T3.2	Preparing future channels (DIERS, VELOMAT, UKT)

The synchronization box ensures that with the start of the recording by the acquisition software all 4 units are sampling images at the same time. The synchronization signal is send also to an external output to give a possibility to synchronize additional signal sources with the video images.

Using the synchronization signal the demands for measuring additional channels are given. This signal can be read in other analysing software tools to measure force, EMG acceleration a.s.o. Due to the fact that especially researchers will work with their own recording and analysing software this is the best solution to ensure maximum flexibility for end users and researchers.

##### Work package T4 (technical): Data organisation and processing

| Work package number: | 4 | Start date or starting event: month 1 |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Participant id: | PW | SMARTTECH | DIERS | KULEUVEN | UKT |
| Person-months per participant: | 36 | 8 | 6 | 2 | 1 |

Objectives:

Processing the synchronised data

#### Task T4.1	Single surface reconstruction	(DIERS, PW, SMARTTECH)

The purpose of single surface reconstruction is to merge point clouds acquired by the four measurement modules into one full 3D point cloud. Dedicated data structures in the time domain and also in frame domain have been developed. Volumetric models based on data of various origins were created. These include data generated by 3D formetric (by Diers) and 3DMADMAC (by Warsaw University of Technology) static measurement systems, as well as data from dynamic measurement system developed for the Aurora project.

Data processing tasks require well considered data structures in order to perform well. Special care must be taken in situations where big amounts of data are processed. The Aurora project is one of these cases.

The experiments have shown that the presented parameters describing surface characteristics are suitable for selecting areas of anatomical landmarks. The parameters are calculated with the use of a non-iterative, numerically stable algorithm. The developed algorithms were implemented and tested on simulated data created with use of computer programs (commercial and developed for Aurora project purposes) as well as on real data acquired with use of 3D and 4D measurement systems. The resulting trajectories are sufficient for animating a simple skeleton model.

#### Task T4.2	Force data and surface parameter processing (PW, DIERS)

It was not possible to perform a test situation where force and kinematic data could be measured synchronously. Both systems are working separately. It is hypothesized that using the synchronization signal the integration of both systems is not a difficult task.

#### Task T4.3	Dynamic reconstruction, morphing, warping	(PW, SMARTTECH)

The proposed dynamic object reconstruction method is based on using an arbitrary mesh which is scaled and wrapped around merged point clouds obtained from measurements, instead of a standard point cloud or standard mesh sequence representation.

The first step is the scanning process of a person who is performing selected movements inside a measurement volume. Then a sequence of merged point clouds is obtained - one for each frame of the measurement. After that an arbitrary 3D mesh is created on the basis of the first point cloud or modelled in a 3D application. Texture UVW mapping is also set at this point. After that for the first data set obtained from the measurement, the arbitrary mesh is scaled and transformed to match the position and size of the point cloud. Then the mesh is wrapped and morphed sequentially around the remaining point clouds. A simplified bone system which helps a wrapped mesh to match a point cloud initially in the following frame is also used. This is necessary in the case of fast movements of a scanned object.

#### Task T4.4	Kinematic object-dependent parameter (KULEUVEN, UKT, PW)

Slices are created using a developed algorithm which is based on a generation of an approximate slice shape from the selected of a point cloud. This is performed by projecting selected points of the point cloud (or mesh) which are in the nearest neighbourhood to the selected slice plane. This neighbourhood is defined by Z which is a height value of the slice plane and the maximum distance ΔZ. During the next step the selected group of points is reduced. This is performed by searching the nearest points to the first point within the specified radius which are deleted afterwards. Next, the nearest point is selected and the algorithm is repeated until the last point is reached. If there are no points which distance is smaller or equal to the specified one (R value) the slice creation is finished. The whole algorithm is repeated for the remaining points so there is a possibility to create several slices. In the case of bad data quality (holes, additional points in a point cloud) there is a possibility of detection of more than two slices. Therefore, the algorithm selects only the biggest two slices. In the case of human bottom body, maximum two slices are detected (one per leg).

Slice area is calculated on using a polygon area calculation method in which selected object is divided into several triangles. The areas of the triangles are added so the final area value is obtained. From the area calculation further clinical relevant parameters like the radius and circumferences can be calculated. In comparison with measuring the circumferences with a flexible meter this method will result in a higher reliability of the measurement. Especially in a repetition measurement the same slice can be detected if an anatomical landmark will serve as a reference. For example if the patella is detected as a landmark the circumference of a defined slice, e.g. slice 15 proximal to the reference, can be compared at different measurement dates.

##### Work package T5 (technical): User interface, visualisation and archiving

| Work package number: | 5 | Start date or starting event: month 3 |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Participant id: | SMARTTECH |  | PW | DIERS |  |
| Person-months per participant: | 10 |  | 10 | 6 |  |

Objectives:

User interface, visualisation and archiving

#### Task T5.1	Task dependent user-friendly interfaces (DIERS)

At this developmental stage no common software surface for grabbing and analysing is available. The acquisition, analysing and visualisation software are at the moment “stand alone” software. They work together, but not under a common interface. This is of course needed for end users. For evaluation and validation and for working under laboratory conditions this point however is neglectable. Compared to the first prototype efforts were made in handling the huge amount of data and in controlling the view from the single cameras. Using switches and different colours for each of the systems a separation of the units is possible. This is especially needed to control the view field of each camera.

Stripes detection:

The task of stripes detection was solved with an 80% - 90% automatic detection and a manually rest-detection and -correction on 10%-20% of the images.

- The user can load one or more directories with the images to be segmented.

- He/she can then select an automatic stripes detection of all the images in the current directory.

- Each image needs 1-2 sec for a complete segmentation and detection.

- The user should do a manually check of the results on each image in difficult cases (weak patterns, false light, etc.).

One challenge was the filtering of background objects and noise. This was done with an object recognition algorithm; sorting out those objects that do not have the characteristic stripes pattern projected on them.

In cases with large homogenous areas, like false light the spots were removed due to the fact that they also missed the characteristic stripe patterns. Also, if the contrast on a pattern-illuminated object is too small (blurred patterns), the object is removed. Small reflection- and shadowed areas cause little problems.

Other phenomena can be observed on the floor and heel region: due to the sensitivity of the stripes detection, stripe pattern that are invisible to the eye are still found on the image. There is a convergence of stripes on the floor and heel where the angle of the projection is big (very narrow distance between the projected stripes). Additional, some noise is falsely detected as stripes. This high sensitivity is necessary, as the illumination may be very weak in some cases. If the algorithm finds such clusters of very narrow segments, they are mostly ignored and not saved as stripes.

Some cases need a manually correction after the automatic detection. This is often the case where the illumination quality is poor: due to weak patters, strong reflections; or if there is a prominent overlap between front and back objects (left thigh and calf is partially hidden below right leg). In most cases this is automatically detected in the segmentation of the image (using the direction of the stripe patterns), but in some cases where factors like bad illumination quality are added to overlap, the algorithms may fail. In those cases, the user needs to manually correct the results by erasing false stripes or segments, or by drawing artificial borders or boundaries between the overlapping objects.

It must be mentioned, that part of the algorithms were developed on the basis of the first prototype. Due to the decision to change cameras the now available images have a much better quality and significant lower noise. This will result in a further improvement of the stripe detection.

#### Task T5.2	Visualisation of models and results (DIERS, SMARTTECH)

Most of the work concerning modeling was done by Smarttech on artificial data. In the last months of the project some visualization was possible with real measurement data and it can be applied for bigger group test. Implemented bone model are simplified and also has to be tested by bigger group of doctors if the result are satisfying and clear enough in most of diagnosis cases. The work performed has been concentrated on landmark tracking of bones which was big task and on collision detection algorithm which is very important when we talk about possibility of showing patient what kind of dysfunction he or she has. This part is probably one of the most crucial parts for successful commercialization of the system.

#### Task T5.3	Compression and archiving (SMARTTECH, PW)

WUT developed rather effective compression algorithms to decrease the amount of data to a acceptable rate. The first step is loading a scan image form the acquired sequence. During the next step the image is filtered to remove noise and get as much details of the stripes as possible. Next the image is being analysed to remove insignificant data by detecting of a foreground (laser stripes in this case) and a background. In this step a significant area mask (which contains detected stripes) is created. Then the background can be separate from the image very easy. Only the information about the lines/stripes is significant concerning the future image analysis. After the above steps have been taken, the image is prepared for a compression. The adaptive method which is based on differential coding using lexical Huffman code is implemented.

In the final compressed file apart from the stripes image data which is not a part of the background a mask table is stored. It is necessary to decompress the image. Mask is created as a binary file using single bits in order to obtain a smaller final file.

The whole process is repeated for each image from the sequence. Finally, after processing the last image all of the compressed data are merged into one archive file. Due to the method described above it is possible to get a high compression ratio without loosing any significant information. The compression ratio of the developed method is about 1:6.17.

The proposed improvements of the compression algorithms seem to be more adequate concerning the data we deal with in the Aurora project. This complex adaptive solution which focuses on an information type provides a very good compression ratio without loosing any significant information.

##### Work package C1 (clinical): Definition of analysis parameters

| Work package number: | 6 | Start date or starting event: month 1 |  |  |
| --- | --- | --- | --- | --- |
| Participant id: | KULEUVEN | DIERS | UKT |  |
| Person-months per participant: | 5 | 3 | 2 |  |

Objectives:

Catalogue of application-dependent landmarks and parameters

Description of work progress:

#### Task C1.1: Anatomical landmark definition (UKT, KULEUVEN, DIERS)

UKT provides the partners with anatomical details and the clinical background. Using the 4D-imager from the 4D-BodyScan Project, good results can be shown in finding landmarks at the knee and ankle joint. The new scan system provides a scan area of 1.2m in height. Therefore restrictions due to the scan field were eliminated. The quality of the landmark detection was a little bit worse with the laser technique than with the light technique. But in principle the defined groups representing an anatomical landmark can be detected. Rather good results were obtained at the knee region. As already shown in previous reports the landmarks coincide rather well with x-ray images and skin markers placed over the bony structure. From the now available data it is obvious that the quality of the images is still improvable. Whilst the system was tested at Smarttech in a laboratory surrounding background noises were rather high. By eliminating this, the quality of the images can be improved.

#### Task C1.2: Definition of parameters related to bony structures (UKT, KULEUVEN)

Axis and angle definition

The definition of anatomical landmarks is based on the convention given by ISB for calculating anatomical models. First results from the DRACO software shows that surface landmarks on the lower leg can be detected. Using the newly developed algorithms from PW it is possible to detect anatomical landmarks automatically and to follow the coordinate over time. In comparison with an electronic goniometer it could be shown, that the knee angle calculation from the DRACO software correlates quite well with the goniometer measurements.

#### Task C1.3: Definition of parameters related to soft tissue (UKT, KULEUVEN)

Besides calculation of axis volume calculation is one very important parameter for clinical diagnosis. Especially after surgery a hypotrophy of a leg is obvious. The developed system allows a direct and objective comparison of the circumference of the affected and non affected leg. Furthermore due to the non-invasive technique, improvements during a rehabilitation process can be quantified.

##### Work package C2 (clinical): Biomechanical modelling

| Work package number: | 7 | Start date or starting event: month 2 |  |
| --- | --- | --- | --- |
| Participant id: | KULEUVEN | PW | UKT |
| Person-months per participant: | 24 | 2 | 1 |

Objectives:

4D anatomical and musculo-skeletal model of the lower limbs

UKT and KUL critically evaluated the prototype system, with regard to a clinical environment. Technical errors were indicated and suggestions were made to improve the robustness of the system. The technical partners will now adjust and improve the system. The next evaluation of the prototype system will include massive recording of a variety of persons doing clinical relevant movements in order to test the developed models and algorithms.

KUL further developed and improved the algorithms for calculation of clinical parameters. All work is done based on a simulated measurement, at a 3D virtual model.

This volume is calculated based on the tracked anatomical landmarks: “centre of rotation of the hip” and “head of femur”. In this way, the change in volume can easily be evaluated during a movement (eg. bending two knees, loaded with bodyweight).

Segmentation of the lower body

A fully automated segmentation was developed based on antropometric data, i.e. the ratio between the lengths of different segments. Initial segmentation is done in an upright posture.  Using the curvature of the contour outline the position of knee joint and ankle can be refined.

The well-know ICP (iterative closest point) algorithm was adapted to track the segments trough a movement sequence. The algorithm uses information about the movement performed in the sequence (movement extrapolation)

Biomechanical model for the hip and knee joint.

Using the relative motion between the pelvic segment and the upper leg segment, the transformation matrices are calculated. With this transformation, the location centre of rotation of the hip joint is estimated. An alternative approach uses a fitted enveloped cylinder as a estimation of the upper leg. Using the relative transformation between the sequential frames, the centre of hip rotation can be calculated.

The model for the knee, developed in the first year of the project is refined. A circle involute model is used to describe the relative movement between the femur and tibia. The form of this circle involute can be derived from the tracked anatomical landmarks. Based on the circle involute model a 4 bar linkage model is calculated. This model is able to predict the movement of the knee joint using forward kinematics. A stick figure model is calculated for each segment in the movement. Using the models desribed above, the accuracy of the models can be improved, also when not enough information is available in the recorded images (for instance when one leg is occluding the other).

Description of the coordinate system

The locale coordinate system for the segments of the lower body is defined using anatomical landmarks. The relation between different segments can be described using the kinematic chain, representing the successive transformation between the segments. The location of the coordinate origin and the orientation of the coordinate system in the world coordinate system will be given for each time step. Also the local coordinates within the coordinate system (as we consider the bones as rigid structures).

##### Work package C3 (clinical): Functional analysis

| Work package number: | 8 | Start date or starting event: month 2 |  |
| --- | --- | --- | --- |
| Participant id: | UKT | KULeuven |  |
| Person-months per participant: | 6 | 6 |  |

Objectives:

Functional analysis

Description of work and progress:

Some preliminary parts of functional analyses were performed using electronic goniometers and also some analysis was done with a VICON system. It could be shown that correlations between these methods and the still existing single view system is possible. Also the calculation of axis could be shown.

Due to the delay in the project however little functional analysis with the working prototype was possible within the official project time. The system is able to generate 3D-point clouds in more or less static situations. The limited sample frequency of 12.5 Hz is to small to analyse functional movements like walking. To achieve this the scanning time of the lasers must be increased. As already mentioned this however will require more energy this is not possible with the selected lasers. Therefore an increase in the frequency may be not possible with the laser technique used.

##### Work package C4 (clinical): Validation and Evaluation

| Work package number: | 9 | Start date or starting event: month 2 |  |
| --- | --- | --- | --- |
| Participant id: | KULeuven | UKT |  |
| Person-months per participant: | 11 | 7 |  |

Objectives:

Validation and evaluation

Description of work:

For validation of the anatomical landmarks, external markers, constructed from plumbum material, were placed onto the skin. Using the DRACO software a direct comparison of the marker placement and the surface image is possible. In a first, primarily qualitative analysis, the correlation between the automatic detected landmarks and the palpated points are acceptable. This analysis however was done only from a single view.

After 21month a 4 unit system was installed in Tübingen. As already mentioned in the 24 month report some problems with the stability of the lasers and the light intensity occurred. Problems could be solved but it was necessary to send the system back to Smarttech for re-calibration. This causes again a delay in testing the system. After reshipping the system to Tübingen measurements from four sides could be performed. Unfortunately the qualities of the images were rather bad and the stripe detection did not result in a satisfying point cloud.

Major reason for the bad quality however was an illumination problem caused by the sensitivity of the cameras from Kamera Werke Dresden. It was tried to solve the problem with additional illumination of the object and to cover the background with black cloth to eliminate background noise. Whilst the background could be kept in black and therefore background noise was minimized the problem of illumination could not be solved. Biggest problem were disturbances caused by shadows and blending effects. Due to this, homogenous stripe detection from four sides was not achievable. Some testing with white tights improved the quality a little bit, but the quality was still not good enough.

Due to this it was decided by the technical partners to change the cameras again. First images results in a much better quality in a dark room. Therefore no additional light sources are needed. The consequence of this decision however was that software adaptations were needed to keep the system running. At the end of the prolongation period real point clouds could be recorded under dynamic conditions.

For example, a person doing a squat exercise or slowly bending one leg while standing on the ipsilateral leg. The 3D reconstruction of these measurements can then be analysed. For each frame the pointcloud can be cut with a plane at several places. Circumference, area and centre of mass can be calculated for each slice. Connecting the centres of mass form successive slices gives the volumetric axis of the focused volume. In this way knee angles can be calculated and compared in different situations. Similar volumes of parts of the leg can be tracked during different exercises.

For UKT however it was too late to finish the necessary clinical evaluation studies. It could be shown from a sphere and as well from a human that a full volume of the lower legs can be analysed with the system. Furthermore from the full volume point cloud transversal cuts can be calculated which represent the radius and the cross sectional area. From the radius also the circumference of the slice can be calculated. The remaining time however was not enough to test the system in some more applications or under different conditions. First real results from the system show, that an objective calculation of the circumference of the lower leg is possible. Compared to normal circumference diagnoses using flexible meters, this technique is more challenging to produce clinical relevant values.

##### Work package O1): Overall Organisation

| Work package number: | 9 | Start date or starting event: month 1 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Participant id: | Diers | R&G | Smarttech | Velomat | PUW | KUL | UKT |
| Person-months per participant: | 6 | 1 | ,5 | 1 | 1 | 1 | ,5 |

Objectives:

Project co-ordination

Description of work and progress:

All the planned regular meetings were arranged by the coordinator. After the kick-off meeting in Wiesbaden, regular meetings were held in Leuven, in Warsaw (PUW), in Kamenz (Velomat) and the midterm meeting in Schlangenbad (Diers). In the second period regular meetings were held in Vilnius (R&G) and two meetings in Tübingen (UKT) Due to the delay of the project another additional meeting was held in Kamenz (Velomat). Besides this, small group meetings were held in Leuven, Warsaw and Tübingen where the prototype was tested.

The final meeting was held in Schlangenbad (Diers).

## Deliverable list

Table 1:Final deliverables list. Due to the delay some deliverables could not be fully realized within the time schedule.

| No | Deliverable name | WP no. | Date due | Actual delivery date | Lead contractor |
| --- | --- | --- | --- | --- | --- |
| D I | A simulated dynamic 3d model | T1 | Week 6 | Week 6 | DIERS |
| D II a | 4D scanner based on split-beam laser components and digital video | T2 | Month 9 | Month 13 | R&G |
| D II b | Force sensor equipment | T2 | Month 9 | Month 13 | VELOMAT |
| D III | Prototype | T3 | Month 12 | Month 16 | DIERS |
| D IV a | Volumetric models | T4 | Month 6 | Month 16 | PW |
| D IV b | Kinetic-Volumetric models | T4 | Month 12 | Month 18 | PW |
| D IV c | Dynamic models | T4 | Month 18 | Month 22 | PW |
| D IV d | Kinematic-Volumetric models | T4 | Month 18 | Month 22 | KULEUVEN |
| D V a | Graphical User Interfaces | T5 | Month 24 | Month 24 | DIERS |
| D V b | Graphical and functional visualisations | T5 | Month 24 | Month 30 | SMARTTECH |
| D V c | Complete archiving and compression structure | T5 | Month 24 | Month 24 | SMARTTECH |
| D VI a | Definition of significant anatomical landmarks in correlation with anatomical structures | C1 | Month 4 | Month 4 | KULEUVEN |
| DVI b | Definition of volume parameters | C1 | Month 6 | Month 6 | KULEUVEN |
| DVI c | Mechanical axis definition and possible correlation with anatomical axis | C1 | Month 6 | Month 6 | UKT |
| DVII | 4D lower body model, including detailed model of knee and ankle | C2 | Month 18 | Month 30 | KULEUVEN |
| DVIII a | Gait and Range of Motion (RoM) analysis | C3 | Month 24 | Following project | UKT |
| DVIII b | Computation of axis in static and dynamic situations | C3 | Month 24 | Month 30 | UKT |
| DVIII c | Data Synchronisation for analysing complex human movement | C3 | Month 24 | Month 30 | KULEUVEN |
| D IX a | Evaluation of clinical relevant parameters under functional conditions. | C4 | Month 24 | Following the project | KULEUVEN |
| D IX b | Establishing of test batteries for functional diagnosis | C4 | Month 24 | Following the project | UKT |
| D X a | Project presentation | O1 | Month 4 | Month 4 | DIERS |
| D X b | Plan for using and disseminating knowledge | O1 | Month 3 | Month 3 | DIERS |

## Milestone List

Table 2. Milestone list with actualized delivery date. Most of the milestones were achieved without the clinical oriented milestones.

| No | Milestone name | WP no. | Date due | Actual delivery date | Lead contractor |
| --- | --- | --- | --- | --- | --- |
| M1 | Initial simulation hardware | T1 | Week 6 | Week 6 | DIERS |
| MII | Dynamic point clouds, scanned within the octagonal cube | T2 | Month 9 | Month 15 | Diers |
| MIII | Synchronized dynamic point clouds with incorporated force data | T3 | Month 12 | Month 18 | VELOMAT |
| M IV | A complete kinetic-kinematic dynamic volume model | T4 | Month 12 | Month 16 | PW |
| M V a | Complete system | T5 | Month 24 | Month 24 | Smarttech |
| M VI b | Catalogue of clinical relevant and detectable/reconstructable parameters | C1 | Month 6 | Month 12 | KUL |
| M VII c | 4D lower body model, including detailed model of knee and ankle | C2 | Month 18 | Month 30 | KUL |
| M VIII a | Functional analysis based on anatomical landmarks, reconstruction of bony structures and soft tissue | C3 | Month 24 | Month 30 | UKT |
| M IX a | Validation and clinical evaluation of automatic landmark detection, as well as the reconstruction of bony structures and soft tissue(ligaments/tendons. | C4 | Month 24 | Following the project | KULEUVEN |

# Section 3 Consortium management

## 3.1. Summarize of the project

A major goal of the project was to develop an analysis system for the diagnosis of lower body extremities using laser stripes from four settings. The composition of the partners therefore requires high competence in the fields of laser technology, image grabbing, movement analysis, biomechanics and clinical/medical knowledge. The composition of the partners seems to be quite well balanced to fulfil the tasks. First technical problems with the laser architecture could be solved by R&G and in cooperation with Velomat, a synchronization unit was developed. The calibration technique was developed by Smarttech and Diers.

Also all the pre-work concerning the clinical and biomechanical oriented task could be done on time and with good results.

However as already mentioned, the status of the project was delayed for about 6 month due to the delay caused by problems in acquiring the cameras. It must be mentioned that the selection of the cameras was the most critical part of the project. Due to the not satisfying image quality the cameras needed to be changed twice. This resulted in a delay because adaptations in the acquisition and stripe detection software were needed. Even when after the last change only minor adaptations in the driver settings of the software were needed, this change caused additional changes in the following processes from acquisition, stripe detection, point cloud generation and full volume calculation. The delayed camera availability therefore resulted in an overall delay of the project working tasks. Some pre-work could be compensated by using the 4D-imager to get point clouds of the lower extremity. A first prototype with one camera and one synchronization unit was constructed by R&G and Velomat and it was shown by them that it worked quite well as a single unit. It was clear that this project deals with a rather new technique and some risks must be taken into account.

The proposed technical task concerning the Hardware is fulfilled just in time. Also most of the Software developments tasks were fulfilled. However the clinical relevant tasks of evaluation and validation the system could not be completed. Real application of the system in a clinical surrounding could not be realized within the project time frame. UKT spent a lot of time in testing the system, with the result, that some corrections were needed. This lead to a further delay, in which no clinical measurements were possible.

As already mentioned to achieve a thorough evaluation of clinical, biomechanical and technical tasks, stable and reliable equipment is needed. The last measurements showed that with the new camera from PGR the system can fulfil this. The clinical evaluation and validation however could not be done and must be postponed to a following period.

## Consortium management task

The former project co-ordinator Mr. Heitmann has withdrawn from the project after 12 month. His work was continued by Mr. Mooshake who is responsible for software development at DIERS. Mr. Heitmann is still working for DIERS as a freelancer and in this way he is still in contact with the project. In this position he provided information for the project coordination especially on all questions concerning stripe detection.

Special task of the coordinator was to invite and organize the regular meetings. All planned meetings could be held at the scheduled meeting table. It might be mentioned here that all meetings were rather successful and well organized.

Table 3: Project timetable and status. Grey cells represents the planned time table for the task Dark is the suggested actual status of the task. Red cells represent task which could not be fulfilled within the proposed time table..

| C3.1 Functional analysis based on anatomical landmarks C3.2Functional analysis based on bony structures T5.3Compression and archiving C1.1 Anatomical landmark definition T3.1Synchronisation of data channels (hardware) T3.2 Preparing future channels (EMG,IR) 0203040506070809101112131415161718192021222324252627282930 T1.1 Prepare and adapt existing equipment T2.1Work Package Bar Chart |  |  |
| --- | --- | --- |
| C4.1Validation of automatic landmark detection C3.3 C1.2Definition of parameters related to bony structures C1.3Definition of parameters related to soft tissue C2.1Tracing of anatomical structures C2.2 Musculo-skeletal modelling C2.34D-modelling C2.4Model interfacing T4.1 Single surface reconstruction T4.2 Force data and surface parameter processing T4.3 Dynamic reconstruction, morphing, warping T4.4 Kinematic object-dependent parameter T5.1Task dependent user-friendly interfaces T5.2Visualisation of models and results Develop optical and laser components T2.2 Integrate optical and laser components T2.3Integrate force sensor components T2.4Implement control and electronic interface software T2.5Implement data channel interfaces T2.6 Calibration 01 |  |  |
| C4.2Functional analysis based on soft tissue reconstruction |  |  |
| C4.3 Validation of reconstruction of bony structures |  |  |
| O1.1 Overall project co-ordination C4.4 Validation of reconstruction of soft tissue/ligaments/tendons | Clinical evaluation |  |

# O1.2 Project presentation                               O1.3 Plan for using and disseminating knowledge                              grey=WP proposed, dark=WP fullfilled, red=WP not fullfilledOther issues

The project does not refer to specific ethical issues.

- Co-operative Research projects

The benefit for the SME’s is that they have direct insight of research results from the RTD partners. Especially for DIERS; as an orthopaedic oriented company the clinical parameters and results can be used directly and indirectly in the medical systems manufactured and distributed by DIERS. The laser company R&G also gets a very useful feedback from the evaluation of their newly developed laser technique. The evaluation of clinical parameters will give information about the practicability of the system for clinical purposes. Furthermore information is available how the laser technology will work in the analysis of human surfaces. This is in most cases rather different from non-living surfaces mostly analysed in technical fields. The same is of course true for Smarttech and Velomat but to a lower degree.

For Velomat it has to be said, that the cooperation with the clinical and biomechanical oriented RTD´s will offer them the opportunity to make a step into the manufacturing of medical products.

Another important issue is the objective evaluation that is being done in universities, providing the SME with reliable sources and references (e.g. articles, presentation on congresses …). This will give the SME’s a higher reputation in this field and improve their positioning in this area.

The SME’s will also profit from new software technology developed at the RTD´s. For instance the video compression and archiving of very large data is one important aspect in this project.

Of course the RTD’s will also benefit from the cooperation with the companies as they get a feeling for economical aspects. Very often discrepancy between economical requirements and science working is obvious. In this cooperation the RTD´s may formulate their “wishes” for this new measurement technology. Due to the fact that the RTD`s have insight in the development process this “wishes” can be specified more in detail and at a time when a changes are possible in the technical development process.

Annex:

- Plan for Using and Dissemination

Exploitable knowledge and its Use

An active dissemination is planned. In today’s world of multimedia and virtual reality, it is of great importance that new methods are presented in a way that the public find “interesting”. The output of this project should be relative easy to “sell” to e.g. television channels and newspapers/magazines. This will create a curiosity and an awareness that will be supported by the solid scientific work done. Also, we hope that many universities/medical clinics will use the system for further studies (thesis, publications, correlation studies,). Even if most of this work will be too scientific for the general public, some papers/demonstrations may be used for public awareness purposes, especially in combination with demonstrations and/or multimedia applications. If the system finds acceptance in general health- or health insurance institutions for e.g. scanning/documentation of population groups, this will also help the awareness.

The eventual use of the technology in non-medical applications like Virtual Reality or the movie industry will greatly benefit the consortium. Also, all kind of industrial applications may benefit from the technology, especially garment industry and sport equipment/clothing industry, and by such applications the general public will for sure get more knowledge and awareness about the technology – which will in turn lead to a greater acceptance of the medical/clinical applications as well.

After further expansion of the 3D formetric system the AURORA product will profit from existing sales channels.

For distribution of the AURORA product different sales channels than the above mentioned may be necessary. It is not possible to predict today the best way of distribution for sales in different countries in a couple of years. Due to the size of the enterprises participating in the project, direct sales with an own organisation seems to be impossible. Therefore the existing dealer network has to be used and has to be expanded.

Another possibility that is being checked is the possibility to realize a global contract with a multinational company with an existing international sales force. In the past, DIERS has had contact with such companies, but due to the wish to stay independent and grow slowly but safely, they have not yet taken this rather drastic step. Nevertheless, this possibility may be necessary for sale and dissemination of the AURORA equipment, and will be discussed and evaluated.

In some countries of Eastern Europe, dissemination will preferably be covered by the project partners SMARTTECH and R&G. Even if they do not have a high medical profile, they know their market structure and are participating in national events, and will help spreading information and knowledge about the product. We expect that, DIERS and SMARTECH as well as R&G will all use the technology gained in the project for expanding their product palette to other related applications.

Table 4.Overview table Exploitable and knowledge.

| Exploitable Knowledge | Exploitable Products | Sectors of application | Timetable for commercial use | Patents/IPR-protection | Owner /Partners involved |
| --- | --- | --- | --- | --- | --- |
| 1. Active Laser Scanning Technology | Active Laser Scanning Equipment | Medical Industrial Multimedia | 2007 |  | R&G (owner) |
| 2. Multiview System for dynamic scanning | Prototype | Medical Industrial Multimedia | 2007 |  | DIERS (owner), SMARTECH (owner), VELOMAT, R&G |
| 3. Dynamic Calibration | Calibrator Master System | Medical Industrial, Multimedia | 2007 |  | DIERS (owner), SMARTTECH (owner) |
| 4. Data Structure of Dynamic Surface | Software / Algorithms | Medical Industrial, Multimedia | 2007 |  | DIERS (owner), SMARTTECH, PW |
| 5. Landmark Detection of Lower Body | Software / Algorithms | Medical, Multimedia | 2007 |  | DIERS (owner) KULEUVEN, PW, UKT |
| 6. Dynamic 360° Reconstruction of MultiView Data | Software / Algorithms | Medical, Research, Multimedia | 2007 |  | DIERS (owner), SMARTTECH, PW (WUT) |
| 7. Reconstruction of anatomical Structures | Software / Algorithms | Medical, Research, Multimedia | 2007 |  | DIERS (owner), SMARTTECH, PW (WUT), KULEUVEN, UKT |

- Dissemination results

1. Active Laser Scanning Technology.

This technology is developed and used in the scanning of dynamic movements. A galvanic mirror laser scanner projects 100-120 horizontal stripes on a moving body with a frequency of 40-60 Hz. With triangulation reconstruction algorithms, it is then possible to reconstruct the scanned surface with a high accuracy.

Main partner involved in this exploitation is R&G. They plan to use the knowledge and technology directly in future products, as well as acting as a contractor for the consortium partners. This technology opens new markets for R&G: Multimedia, Medical and Industrial dynamic reconstruction.

2. Multiview System for dynamic scanning.

The synchronized Multiview hardware is developed by R&G, VELOMAT and DIERS. 4 Laser-Camera Stereo pairs are working together with other external sensors (pressure, etc.) in order to scan as many biometric data as possible of a movement sequence. The collected data will then be used as a basis for anatomical, biomedical and medical analysis. The possibilities are vast, if the reconstruction works: The ability to follow a detected anatomical structure with a high surface accuracy, gives the user the possibility to analyse clinical parameters to a very high degree. Also, the equipment can be user for scanning of any moving object (with or without structural changes), which opens the possibility to use the system in industry, multimedia, etc. As far as the consortium knows, currently there exists no comparable technology: low-cost, easy to use, high accuracy, flexible.

3. Dynamic Calibration.

The methods and algorithms of the two-step calibration have been developed by SMARTTECH and DIERS. Step 1 calibrates the 4 single-sided scanners, and Step 2 combines the views. This calibration is rather universal, and may be used for any dynamic or static 3d scanner. Thus, the possibilities for using the methods in other equipments as well in research projects (easy to understand and adapt) makes it very suitable for reuse for both SME and RTD performers. The consortium expects this to be the basis in future products, projects and studies.

4. Data Structure of Dynamic Surface.

The data structure of the huge amount of data collected and scanned is defined as a sort of moving topology. The exploitation of this topology will without doubt be very interesting both in industrial commercial products as well as in research institutions. PW (WUT) has developed this data structure, and they will continue the research as there are some very big advantages by using this topology; especially in viewing (very easy interface to animations) and storing (relative few nodes compared to the original data). All partners may use this topology in any future and current technology, thus it will be a basic feature in products.

5. Landmark Detection of Lower Body.

Building on the results of a former project (4D BodyScan), the RTD performers and DIERS are extending anatomical landmark detection to lower body. This development has many applications, as it opens up the possibility of following movements of anatomical structures (especially bones) below the surface. All further commercial products build on this ability, and it gives the SMEs a big advantage compared to competitors, as competing systems (i) do not reconstruct complete surfaces, only points and/or (ii) do not automatic detect the structures but need time-consuming and advanced operator-interaction (markers, post-segmentation, etc).

6. Dynamic 360° Reconstruction of MultiView Data

An accurate, robust and dynamic reconstruction of a complete lower body including kinetic and kinematics data is the basis for clinical and medical applications. The novelty of this equipment gives the partners a big advantage in further development of (i) clinical (ii) multimedia and (iii) industrial applications. Developed by PW (WUT), DIERS and SMARTTECH; they will be able to insert the methodology immediately in commercial products. The Multiview Equipment for dynamic scanning and the dynamic calibration results are defining the basis for this knowledge.

7. Reconstruction of anatomical Structures.

Developed by the RTD partners, this reconstruction makes the system user friendly and “sellable“. Some very new results will be the outcome of a movement study with a reconstruction of anatomical structures. This will give especially DIERS a leading edge in medical computerized diagnostics, but also the RTDs will benefit from this as the possibility to publish the results is very good.

- Dissemination strategies

Because the exploitation of the AURORA system and the launching of the product will take place as soon as possible after finalising the developments, the dissemination strategy is part of the project to expand the new technology to other sectors. Other sectors could be aesthetic surgery, shoe industry, garment industry etc. Also other companies with the demand of contact-free dynamical measurements could benefit from the development. Finally, the project results may contribute to the European technology progress in general, due to the novelty in 4D measurements of the lower limbs in combination with lower costs.

The consortium expects the opportunity to disseminate the project results using the channel of national relay centres. PR activities, publications, magazines, TV etc. are other tools to disseminate the results of the development.

The dissemination of the product resulting from the AURORA project has to be divided into periods.

- The first period was the time of development (as the project ran). During this time, the participants, especially the RTD-performers, presented the scientific results and application possibilities on congresses, trade fairs etc.

- At the end of the project period the product launching will start. The launching and distribution of AURORA will be based on existing and new distribution channels: those already existing for the product formetric (3D-back- and spine measurement, see below), as well as the new dissemination channels for the 4D BodyScan equipment.

- Additional, after and during introduction of the product to market, supplementary worldwide distributors have to be found. The potential is big, as we have noticed for the formetric system, especially on congresses and trade fairs: potential distributors are eager to bring the products to their assortment.

- Distribution:

Today DIERS is manufacturing and distributing the 3D and 4D back and spine system formetric in several countries. In Germany, DIERS is controlling the whole manufacturing, distribution, sales and service network. In other European countries sales takes place both with local distributors, as well as direct distribution to certain degree. In Italy the formetric system is strongly presented, thus having the strongest distribution channel in Europe (beside Germany). Systems are additionally installed in Egypt, South Korea, Singapore and Malaysia. Negotiations with a distributor for Canada and the US are in progress, and a market introduction is expected very soon.

- Dissemination of Knowledge

DIERS is participating in the most important German congresses and is exhibiting the 3D formetric on several international fairs:

- The most important event for international business and to create contacts is the MEDICA fair taking place every November in Düsseldorf/ Germany. For many years DIERS is taking part with an own booth.

- The first time in 2003 DIERS was an exhibitor on the American Congress of Orthopaedic and Orthopaedic Surgeons.

- DIERS is participating in the world congress of Orthopaedic Technicians (OT) taking place every 2 years in Leipzig/ Germany.

In addition, DIERS participates annually in the two main national orthopaedic congresses DGOOC (German Association of Orthopaedics and Orthopaedic Surgery, Berlin) and the South German Congress of Orthopaedics (Baden Baden). All those activities are leading to more international distribution channels and will be used later for the dissemination.

The outcome of this project will be integrated into the courses offered by the universities involved, K.U.Leuven, University of Tübingen, Warsaw University of Technology. Through their participation in international student exchange networks, some of them sponsored by the EC, dissemination will be wider spread than by the universities themselves.

- Publications/presentations:

International and national conferences, congresses, exhibition and trade fairs are excellent media to reach both interested customers as well as possible distributors. An active participation on such events was understood for both the SMEs and for the RTDs. There was an active support from the side of the SMEs regarding the participation of the RTDs on conferences and congresses, as well as publications and other scientific work (within the framework of protection of knowledge, as specified in the contract/agreement and their annexes). In the following table congress presentations from the RTD partners are listed.

Table 5: Congress presentations in 2005 and 2006 concerning the project.

| Congress venue | presenter | Title of the presentation | year |
| --- | --- | --- | --- |
| SPIE International Congress on Optics and Optoelectronics Warszawa, Poland | WUT | Automatic anatomical structures location based on dynamic shape measurement. | 2005 |
| 20th Congress of the International Commission for Optics 21-26 Changhum, China | WUT | Opto-numerical procedures supporting dynamic lower limbs monitoring and their medical diagnosis. | 2005 |
| XXth Congress of the International Society of Biomechanics Cleveland, US | UKT | Calculating anatomical leg structures from surface contours | 2005 |
| IFMBE, Prague, Czech Republic | KUL | A dynamic contact-free optical acquisition technique for anatomical structures of the trunk and lower body. | 2005 |
| Kongress der Deutschen Gesellschaft für Sportmedizin (DGSP) Hamburg, Germany | UKT | Erfassung anatomischer Strukturen der unteren Extremitäten anhand von Oberflächenanalysen | 2005 |
| Electronic Imaging, San Jose, US | WUT | 4D data processing for dynamic human body analysis | 2006 |
| Photonics Europe, Strasbourg, France | WUT | 4D measurement system for automatic location of anatomical structures, | 2006 |
| 9th Symposium on 3D Analysis of Human Movement Valenciennes, Spain | WUT | 4D measurement and surface analysis for diagnosis of human lower limbs motion | 2006 |
| Triennial Conference of the International Ergonomics Society. Maastricht, The Netherlands | KUL | A dynamic contact free optical acquisition technique for trunk and lower body | 2006 |
| World congress of Biomechanics Munich, Germany | UKT | Evaluation of newly developed algorithms for calculating anatomical landmarks from surface contours on the lower extremities. | 2006 |
| Schmerzsymposium Nottwill. Nottwill, Schweiz | UKT | Möglichkeiten der Oberflächenanalyse der unteren Extremitäten mittels Musterprojektion | 2006 |

- Design study:

Following a design study a prototype was executed. To the system finally brought to market, some more inspiration will be gathered in order to achieve an attractive system. The prototype available the cameras and laser scanners were mounted into a fixed frame which is mounted on an elevator. This will increase the application possibilities of the system because the measuring volume can be changed. In combination with the improved calibration process a higher flexibility of the system can be achieved.

- Follow up studies

One very important part of the acceptance of the AURORA system is to create a plausible and solid clinical validated “backbone”; in the meaning that the consortium needs to go to leading academic/research institutions in order to actively support the commencement of studies, publications, thesis and papers. The co-ordinator DIERS already has a very broad and close co-operation with leading universities and medical clinics – especially in Germany. The RTD-partners are all well-known and respected institutions, world-wide acknowledged for their excellent scientific work. This creates a unique possibility to initiate irreproachable follow-up studies for scientific/clinical validation from neutral institutions, and the consortium will be open and supportive for such possibilities.

- Non-scientific applications:

The possibilities for the equipment to be used in non-scientific applications are existent, and will be examined thoroughly. Especially for the use of the system in all kind of multimedia activities, there is a great potential (virtual reality, movie industry, leisure activities). In any case, there is a big opportunity to present the results in an “exiting” form, with animations and public-friendly interfaces, an opportunity that also will create more awareness about the product in the general audience.